CHIRALYSER - MP

Instruction Manual

Version CHI 7.3x   (LED 426 nm)

Table of contents

1.0      Notes and Warnings

  - Function Principle

- Installation

2.1      Electrical connections

- Start of Measurement

3.1       Peak detection

- Change Instrument settings

4.1       RANGE

4.2       AVR (time constant)

4.3       BASE (baseline offset)

4.4       NOISE

4.5       ASC (Automatic Symmetry Calibration)

5.0      Troubleshooting

6.0      Service Instructions

6.1      Disassembling the Flow Cell

6.2      Mechanical Adjustment of the Optical Balance

- Preparative  Flow Cell

8.0      Pin connection

8.1      Remote control via serial Interface

IBZ MESSTECHNIK GMBH                       Telefon  05041 / 757877

Buchenkamp 3                                          Telefax   05041 / 758213

D – 31832  Springe                                    info@ibzmesstechnik.de

## Germany                                              www.ibzmesstechnik.de

1.0  Notes and Warnings

This apparatus has been designed and manufactured in accordance

with EN 60335/1 resp. IEC / CE and has left the factory in perfect

condition as regards the relevant safety and CE regulations.

In order to maintain this perfect condition and to ensure safe

operation, the user should observe the notes and warnings

contained herein.

Prior to switching this apparatus on, please make sure that the

mains voltage corresponds to the voltage stated on the apparatus

and that the mains supply is correctly grounded. The protective

effect must not be reduced by any unearthed extension lead.

Any interruption of the earth wire inside or outside the apparatus

or loosening of the earth wire connection can make an apparatus

dangerous. Intentional interruption is not permissible.

Opening of covers or removing of parts may expose live parts.

Connections points may also be live !

Before making any adjustments, servicing, overhauls or before

replacing any parts, the apparatus must be disconnected from

the mains supply, if it needs to be opened.

If it becomes necessary to work on the apparatus whilst it is

open, this must only be carried out by a skilled person, who

is familiar with the dangers involved.

Care has to be taken to ensure that only fuses of the indicated

type and nominal amperage are used. It is neither permissible to

use mended fuses nor to short-circuit the fuse holder. The

apparatus should not be operated, if there are any doubts about

its safety, and it should then be taken out of service and

secured against unintentional use.

Reasons for suspecting the apparatus to be unsafe, include the

following:

---  visible damage,

---  no function,

---  prolonged storage under unfavourable conditions,

---  severe mechanical stress due to unsuited transport.

1.1   Function Principle

From the LED (1) light  passes  through the lens system (2) and gets

linear polarized by the polarization filter P1 (3).

After that, the light passes the flow cell (4) and the Faraday coil (5) with

the Faraday rod (6) inside.

The Faraday rod consist  of a special  glass  material which rotates the

plane of the  polarized light  synchronous  to the frequency f1 of the

electrical field applied to the coil. (Faraday modulation).

After passing the second polarization filter P2 (7), which is cross

arranged to the first polarization filter, the light is modulated twice

the frequency f1.

The voltage controlled photomultiplier (8) converts the modulated light

into a photocurrent.

In the electronic synchronisation stage (10) the phase of the photocurrent

is compared with the phase of the modulator frequency f1.

An optical active compound in the  flow cell rotates the plane of the

polarized  light and causes a phase-offset in the synchronisation stage.

This phase-offset generates via the power amplifier (11) a compensation

current  into the Faraday coil which causes an electrical field which

compensates the rotation of the optical active compound.

(Faraday compensation).

The compensation  current is  proportional to the  optical rotation of  the

optical active compound. Positive (d) - or negative (l) rotations are

are displayed (12) by positive or negative currents.

2.  Installation

Place the CHIRALYSER nearby to the outlet of the column or to

the outlet of any detector on a solid and vibration-free location.

There must be no other objects placed upon the instrument.

Connect  the IN port  with a 1/16" capillar with an inner diameter

of  0.25 mm (or less) and suitable fitting to the outlet of the column

or  the outlet of a previous detector.

Connect the OUT port with a 1/16” capillar with an inner diameter

of  0.7 mm (or more) and suitable fitting to the system waste.

The CHIRALYSER has to be always the last detector connected

to the HPLC system because of it`s sensitivity to back pressure

caused  by other detectors.

Disregard will have high risc for braking the thin glass windows

of the flow cell.

Attention:

There must be no back pressure  at the OUT port.

No else detectors may connect to the OUT port.

# Only use degassed mobile phase and do not exceed

maximum flow rate ( 3 ml/ min for analytical cell).

2.1 Electrical connections

Mains

Make sure that your ac mains voltage is identical to the value printed

on the identification label at the back side of the instrument.

Operation without  proper ground connection is dangerous and will cause

malefunction of the instrument.

Analog Interface and AUTOZERO

Connect  the analog cable to your system`s A/D converter input port or

to an Integrator or Chart Recorder. Notice the polarity marked on the cable.

The output voltage range of the CHIRALYSER is 1000 mV related to the

selected  baseline offset and measuring range.

Default setting is 500 mV.

The analog cable also carries the signal for the AUTOZERO contact.

Contact has to been closed for minimum 0.5 sec to create ZERO in

the display. There must be no external potential (voltage) at the

contact.

IMPORTANT : Use only original analog cable attached to the consignment.

See section 8.0

Serial Interface

Connect  the serial cable to your system`s RS232c input port.

Configurate your system`s serial input port to:

9800 baud, 8 data,1 stop, 1 start, no parity.

The actual measuring value is obtained by sending “?” (ASCII  63)

to the CHIRALYSER.

IMPORTANT : Use only original serial cable attached to the consignment.

For full remote control via serial interface see section 8.1

USB Interface   (optional )

Connect the USB adapter to the serial interface connector and to your

system`s USB port .  Start the driver CD and follow the instructions

given by the installation program.

Analog interface overrange.      Touch  ZERO

Peak detection, see  section 3.1

Automatic Symmetry Calibration, see section 4.5

3.0   Start of measurement

After switching POWER on, the CHIRALYSER displays

Wait until flow and pressure  is balanced and stable and then

touch   START.

Flashing    α    indicates correct data sampling.

CHIRALYSER is ready for measuring.

Touch  ZERO  to start  operation in the default mode.

Display may show some additional  information.

( see Fig. 5 – 7, left side ).

  - Peak detection

An optically active compound is indicated by  peak in the display.

Baseline drift and noise, specially during warm up (approx.30 min.),

will also be indicated by peak.

Increase time constant (AVR)  if baseline noise and drift

- under balanced conditions -  will indicate permanently PEAK

at the display.

See section 4.0

4.0    Change instrument settings

Touch  MENUE  for changing the instrument settings

CHIRALYSER will display

Select parameters to change settings.

Touch  BACK to go back to the measuring mode.

4.1   RANGE

Default setting for the measuring range is 8.0 mdeg.  / 1000 mV full scale.

Range limits are 4.0 mdeg. / 1000 mV  for highest sensitivity and

512.0 mdeg.  / 1000 mV  for lowest sensitivity.

Select appropriate RANGE by touching the arrow keys.

Touch OK to go back to the measuring mode.

  - AVR  (time constant)

Default setting for the AVR time constant is “OFF” (no time constant).

In this setting baseline noise is visible, but lowest limit of detection

is attainable.

AVR settings can be changed within 0.2 sec (fast response) to 3.6 sec

(slow response).

The AVR settings may affect baseline noise and peak geometry.

Select appropriate AVR by touching the arrow keys.

Touch OK to go back to the measuring mode.

4.3  BASE

Default setting for the baseline offset is 500.0 mV.

Baseline offset limits are 20.0 and 980.0 mV.

Use default setting for symmetrically measuring range.

Select appropriate BASE by touching the arrow keys.

Touch OK to go back to the measuring mode.

  - NOISE

The displayed noise value effects the dynamic discrimination

level of the AVR settings. Readings are only for diagnostic

purposes and cannot be changed by user.

Touch OK to go back to the measuring mode.

  - ASC  ( Automatic Symmetry Calibration )

Internal optical offset, caused by high refractive index of the mobile

phase or mechanical stress to the cell windows will display flashing

ASC in the operational mode (Fig. 7).

Touch MENUE  (Fig. 8) and then ASC to start the

Automatic Symmetry  Calibration.

CHIRALYSER will display

When terminated, CHIRALYSER will display

Touch  OK  to go back to the measuring mode.

5.0   Troubleshooting

Error message        NO SIGNAL

Meaning                  Air bubbles inside flow cell.

Windows unclean or broken.

Remedy                   Use only degassed mobile phase.

Raise flow temporary to max. 3 ml/min.

Purge with isopropanol to remove bubbles.

Clean or replace windows. ( see section 6.1)

Error message        ASC OVERRANGE

Meaning                  Internal optical offset is  too high,

Automatic Symmetry Calibration cannot

Correct the optical offset.

Remedy                  Adjust optical balance ( see section 6.1)

Error message        Display indicates PEAK at  baseline

Operation.

Meaning                   Baseline is noisy.

Small air bubbles inside flow cell.

Instrument  not yet thermically balanced.

Incompatible mobile phase

( e.g.: organic / water).

Remedy                  Wait until instrument is balanced.

Inject isopropanol to remove bubbles.

Use compatible degassed mobile phase.

Error message        Baseline drift

Meaning                   Instrument not yet thermically balanced

Remedy                   Wait until instrument is balanced.

Keep temperature constant.

6.0     Service instructions

This chapter contains some easy to do service instructions.

The CHIRALYSER supports the user with diagnostic help.

Opening of enclosures or removing  covers may expose live parts.

Connection points may also be live !

If it becomes  necessary to work on the apparatus whilst it is

open,  this must only be carried out by a skilled person, who is

familiar  with the dangers involved.

See: Notes and warnings at the beginning of this manual.

For the following service instructions, the enclosure has to be open.

Switch off power and open / remove the screws according this drawing.

Bend backplane slightly and lift off the upper part of the enclosure.

6.1     Disassembling the flow cell

The flow cell must be disassembled if the windows are broken or

require cleaning.

For do so, follow these instructions:

-    Switch off  power  and  open the enclosure according the drawing

in section 6.0

-    Disconnect the 1/16" fittings at the  flow cell (4).

-    Remove the cell holder (clamp); take the flow cell to a clean

place for disassembly.

-    Open the flow cell by  unscrewing  the screw plate (1) with the

assembly tool (5).

-    Replace the defective or unclean window  together with its VITON

seal (2) The  area around and under  the window must be absolutely

clean; wipe away any particle very carefully.

-    Carefully screw the flow cell closed with the assembly tool with very

light torque.

-    Repeat the same procedure  with the  other side of the flow cell.

-    Very carefully, tighten both sides – slowly  with little torque.

If a screw plate is tightened too quickly, or with too much torque,

stress may be induced to the windows, causing baseline drift.

-    Fix  the cell with the cell holder (clamp) and reconnect the 1/16“ fittings

to the flow cell.

-    close the enclosure and fix screws.

Switch the instrument on and touch START.

Expect some  drift  for some time, until stress is distributed

in the mechanical / optical system and the instrument gets

balanced.

6.2   Mechanical Adjustment of the optical balance

The CHIRALYSER is  factory set  to an symmetrically  measuring  range

with  excellent  optical balance.

This balance may be upset through flow cell disassembly or mechanical

shock  by transport  and may cause peak cut off and internal optical offset.

This internal optical offset is indicated by displaying a flashing ASC in the

operational mode.

Starting the ASC routine (see section 4.5), will cancel the optical offset.

In case of excessive optical offset, indicated by ASC overrange (see Fig. 16),

the  Automatic Symmetry Calibration cannot cancel the optical offset

completely and needs mechanical adjustment of the optical balance.

During the mechanical adjustment, power is on and flow rate should be

1 ml/min.

Do not touch any electronic device nor  the printboard inside

the instrument. There is a high risk of electric shock.

Open the enclosure according the drawing in section 6.0

Mark the position of the symmetry screw with a soft pencil.

Slowly rotate the symmetry screw ( Fig. 17) clockwise until  display

indicates “please wait” (Fig. 13).

Do not rotate more than 5 turns clockwise. If there is no result, rotate

anti-clockwise for maximum 10 turns.

Close enclosure and continue operation.

7.0  Preparative Cell

##### This CHIRALYSER is optionally provided with a preparative measuring cell

with an optical length of 100 mm, volume 300 µl.

Recommended flow rate is 5 to 100 ml/minute.

Do not exceed 30 bar (450 psi) pressure at the outlet (OUT) port

and use only degassed mobile phase specially when flow rate is low.

The detector may be much too sensitive in its default settings.

To reduce sensitivity, first make an ASC (Automatic Symmetry Calibration)

with suitable flow according section 4.5 in this manual and then select

appropriate RANGE setting.

If sensitivity is too high even in the highest RANGE setting (512.0 mdeg.),

a shorter measuring cell is recommended.

###### Installation

##### The 100 mm cell is placed  towards the light source and the “IN”

tubings must be changed to 1/16” with inner diameter 1.0 mm to

minimize backpressure to the column.

Fix clamp midside of the cell.

8.0   Pin connections

The Analog cable carries the signal (1000 mV) for the external A/D converter

and  the connection for the AUTOZERO contact.

Analog cable    ( 5 pole connector  male / open )

Pin                               meaning                                               color

1               Signal Ground                                                       green

2               Auozero  - close contact to pin 1 -                        grey

3

4               analog signal  negative potential                         white

5               analog signal  positive  potential                           blue

Serial interface cable      ( 9 pole D-sub connector   male / female )

Pin                     CHIRALYSER (female)

1                          DCD

2                          TxD       (transmit data)

3                          RxD       (receive data )

4                          DTR

5                          GND      (signal ground)

6                          DSR

7                          RTS

8                          CTS

9                           n.c.

  - Remote control via serial interface

##### Remote control is given by sending ASCII code to CHIRALYSER according schedule

##### FORMAT  9800 / 8 / 1 / 1

##### ASCII                          mode                                                action

65         (A)               START                                   start CHIRALYSER

66         (B)                ZERO                                    set display ZERO

87         (W)             DISPLAY                                        display off

75         (K)              DISPLAY                                        display on

63         (?)       REQUEST DATA                  send measuring data to host

67         (C)            MENUE                            enable MENUE for all settings

73         (I)              MENUE                               back to measuring mode

68         (D)             RANGE                                   enable RANGE settings

76         (L)              RANGE                                       increase RANGE

77         (M)             RANGE                                       decrease RANGE

78         (N)             RANGE                              back to measuring mode

70         (F)                 AVR                                          enable AVR settings

82         (R)                 AVR                                               increase AVR

83         (S)                 AVR                                               decrease AVR

84         (T)                 AVR                                  back to measuring mode

72         (H)              BASE                                        enable BASE settings

79         (O)              BASE                                           increase BASE

80         (P)              BASE                                           decrease BASE

81         (Q)              BASE                                    back to measuring mode

69         (E)                HV                                                 enable HV

85         (U)                HV                                       back to measuring mode

71         (G)               ASC                                                enable ASC

86         (V)               ASC                                      back to measuring mode